EMBO

Embolization Particle Sizing System

Custom MCU Board — PCB Design Brief

| Field | Detail |
| --- | --- |
| Document version | v2.51 — Errata release: ESP32 module single 3V3 pin correction (§2.2); AD9833 RSET/COMP/200Ω pin errors corrected (§5.1) |
| Date | June 2026 |
| Status | For schematic design |
| MCU module | ESP32-S3-WROOM-1-N8 (8MB flash, PCB antenna, no PSRAM) |
| Motor drivers | 2× BigTreeTech TMC2209 V1.x plug-in modules, socketed |
| UI system | Breakout board via 20-pin IDC ribbon: ILI9341 TFT, EC11 encoder, 2 buttons, passive piezo buzzer |
| Reverse polarity | AO4407A SOIC-8 P-MOSFET + BZX84C15 15V Zener gate clamp |

v2.51 ERRATA — two datasheet errors corrected that have been present since v2.1: (1) Section 2.2: ESP32-S3-WROOM-1-N8 module exposes exactly one external 3V3 supply pin (castellated pin 2). VDD3P3, VDD3P3_RTC, VDD3P3_CPU are internal die rails, not external pads. Decoupling table reduced from 3 rows to 1 row. 10µF bulk cap count reduced from 5 to 4. 100nF count reduced from 25 to 23. (2) Section 5.1: AD9833 MSOP-10 has no RSET pin (that pin belongs to the AD9834). The 6.8kΩ 1% precision resistor and 604mVpp RSET-based amplitude calculation are removed. The pin previously called "COMP" is correctly named CAP/2.5V (pin 3) — 100nF to DGND. The 200Ω external load resistor is removed — the AD9833 VOUT has an internal 200Ω output impedance; no external load resistor is required per Rev G datasheet. Output amplitude is approximately 0.5Vpp at 3.3V supply, fixed internally. The Tx amp Rf=390Ω choice remains valid and safe with this corrected input amplitude.

# Revision history

| Version | Summary of change |
| --- | --- |
| v2.0 | STM32G474 → ESP32-S3-WROOM-1-N8. USB-C native programming. BLE debug. |
| v2.1 | UAS gain correction. TMC2209 VCP + SPREAD fixes. Power hardening. |
| v2.2 | AO3401A → AO4407A + Zener. Motor 4×2-pin → 2×6-pin. KiCad footprints. |
| v2.3 | ILI9341 confirmed, 14-pin header. RPi UART JST-XH 3-pin. C_adc explicit. USB-C pins documented. |
| v2.4 | UI system moved to breakout board via 20-pin IDC ribbon. |
| v2.5 | TMC2209 QFN-28 → BTT module sockets. SpreadCycle firmware-enforced. CLK pin correction. |
| v2.51 | ERRATA: ESP32 module decoupling corrected (single 3V3 pin). AD9833 RSET/COMP/200Ω errors corrected (§2.2, §5.1). |

# 1. System overview

EMBO is a closed-loop embolic agent preparation device. The MCU board integrates three sensor subsystems (UAS, CV camera, StallGuard viscosity), two stepper motor actuators, two mechanical limit switches, and a UI breakout board connected via IDC ribbon cable.

## 1.1 Why ESP32-S3-WROOM-1-N8

- Module: chip, 8MB flash, 40MHz crystal, PCB antenna — no external MCU support components needed

- Built-in USB Serial/JTAG on GPIO19/20 — no bridge IC needed

- Bluetooth LE 5.0 on-chip — NimBLE for wireless debug

- 36 GPIO broken out — 28 used in this design, 8 free. N8 variant: GPIO35–37 free for SPI bus.

## 1.2 Power rails

| Rail | Source | Voltage | Consumers | Max current |
| --- | --- | --- | --- | --- |
| 12/24V | PSU direct | 12–24V | 2× TMC2209 module VM | 6A peak |
| 5V | LM2596 buck | 5V ±2% | LED panel, TMC2209 module VCC_IO, buzzer | 2A |
| 3.3V | AMS1117-3.3 LDO | 3.3V ±1% | ESP32, AD9833, IC1/IC2, TFT logic | 800mA |

# 2. ESP32-S3-WROOM-1-N8 module

## 2.1 Footprint and keepout

- Dimensions: 18.0×25.5×3.1mm. Castellated pads, 1.27mm pitch.

- ANTENNA KEEPOUT: no copper, traces, vias, or components within 15mm of PCB antenna on ALL layers.

- EPAD: solder to GND pour.

## 2.2 Power decoupling — CORRECTED in v2.51

CORRECTION from v2.1–v2.5: prior versions listed VDD3P3 (×2) and VDD3P3_RTC as separate decoupling rows, implying three external supply pads. The ESP32-S3-WROOM-1-N8 module exposes exactly one external 3.3V supply pin — castellated pin 2, labeled 3V3. VDD3P3, VDD3P3_RTC, and VDD3P3_CPU are internal die rails routed inside the module. The decoupling table below reflects the actual single external pin.

| Pin | Net | Decoupling |
| --- | --- | --- |
| 3V3 (castellated pin 2 — only external 3.3V supply pin) | 3.3V | 10µF X5R 0805 + 100nF X7R 0402 — place within 2mm of this pad |

## 2.3 EN reset + boot circuits

| Component | Value | Connection |
| --- | --- | --- |
| R_en | 10kΩ 0402 | 3.3V → EN pin |
| C_en | 1µF X7R 0402 | EN → GND |
| SW_reset | 6×6mm tactile | EN → GND |
| R_boot | 10kΩ 0402 | 3.3V → GPIO0 |
| C_boot | 100nF X7R 0402 | GPIO0 → GND |
| SW_boot | 6×6mm tactile | GPIO0 → GND |

## 2.4 USB-C programming connector (J13)

| USB-C signal | PCB connection | Notes |
| --- | --- | --- |
| D− A-side (A7) | GPIO19 | Tie A7+B7 on D− net → GPIO19. |
| D− B-side (B7) | GPIO19 (same net) | A7+B7 shorted on PCB. |
| D+ A-side (A6) | GPIO20 | Tie A6+B6 on D+ net → GPIO20. |
| D+ B-side (B6) | GPIO20 (same net) | A6+B6 shorted on PCB. |
| CC1 (A5) | 5.1kΩ to GND | Mandatory. |
| CC2 (B5) | 5.1kΩ to GND | Mandatory. |
| VBUS | 5V via 1N5819 |  |
| GND | GND |  |
| SBU1 (A8) | NC |  |
| SBU2 (B8) | NC |  |
| SSTX/SSRX | NC |  |

# 3. Power section

## 3.1 Reverse polarity — AO4407A + BZX84C15 Zener

| Component | Value | Connection |
| --- | --- | --- |
| Q1 | AO4407A SOIC-8 P-ch −30V −12A | Source (2–4) → J14 V+. Drain (5–8) → 24V rail. Gate (1) → R_gate/D7 junction. |
| R_gate | 10kΩ 0402 | Gate → GND |
| D7 | BZX84C15 15V Zener SOT-23 | Cathode → Source. Anode → Gate. Clamps Vgs=−15V. |

## 3.2 Input protection

| Component | Value | Connection |
| --- | --- | --- |
| F1 | 1A polyfuse (1812) | In series after Q1 |
| D5 | SMBJ28A TVS | Across 24V rail after F1 |
| Power LED | Red LED + 10kΩ | 3.3V → 10kΩ → LED → GND |

## 3.3 Buck: 24V → 5V (LM2596-5.0)

| Parameter | Requirement |
| --- | --- |
| IC | LM2596-5.0 fixed output |
| Input | 8–40V |
| Output | 5V 3A |
| Catch diode | SS34A — within 5mm of SW pin. Snubber: 10Ω + 1nF. |
| Bleed resistor | 4.7kΩ from 5V to GND |

## 3.4 LDO: 5V → 3.3V (AMS1117-3.3)

| Parameter | Requirement |
| --- | --- |
| IC | AMS1117-3.3 SOT-223 |
| Output | 3.3V 800mA |
| Thermal | SOT-223 tab → ≥100mm² copper + ≥4 thermal vias. ≈850mW at full load. |

# 4. GPIO pin assignments

SPI mode: AD9833 Mode 2 (CPOL=1). ILI9341/XPT2046 Mode 0. XPT2046 max 2MHz. ILI9341 via ribbon: 20MHz.

| GPIO | Signal | Dir | Connected to | Notes |
| --- | --- | --- | --- | --- |
| 1 | UAS_ADC | Analog in | Envelope output | ADC1_CH0. 100nF C_adc within 2mm. |
| 2 | STATUS_LED | Out | 330Ω → LED → GND |  |
| 4 | TMC_UART | Bidir | TMC2209 PDN junction | See §6.3 |
| 5 | STEP_M1 | Out | Module U5 STEP | LEDC PWM |
| 6 | DIR_M1 | Out | Module U5 DIR |  |
| 7 | EN_M1 | Out | Module U5 EN | 10kΩ pull-up on main board |
| 8 | STEP_M2 | Out | Module U6 STEP | LEDC PWM |
| 9 | DIR_M2 | Out | Module U6 DIR |  |
| 10 | EN_M2 | Out | Module U6 EN | 10kΩ pull-up on main board |
| 11 | BTN1 | In | IDC pin 16 | Breakout pull-up + debounce |
| 12 | BTN2 | In | IDC pin 17 | Breakout pull-up + debounce |
| 13 | BUZ_PWM | Out | IDC pin 18 | LEDC PWM |
| 14 | LIMIT_M1 | In | Limit switch (J6) | 100nF debounce |
| 15 | LIMIT_M2 | In | Limit switch (J7) | 100nF debounce |
| 16 | EC11_A | In | IDC pin 13 | Breakout pull-up + debounce |
| 17 | EC11_B | In | IDC pin 14 | Breakout pull-up + debounce |
| 18 | EC11_SW | In | IDC pin 15 | Breakout pull-up + debounce |
| 19 | USB D− | Fixed | J13 A7+B7 | Fixed USB PHY |
| 20 | USB D+ | Fixed | J13 A6+B6 | Fixed USB PHY |
| 21 | TOUCH_IRQ | In | IDC pin 12 | 100nF near J9 |
| 35 | SPI_MOSI | Out | AD9833 + IDC pin 5 | SPI2 |
| 36 | SPI_CLK | Out | AD9833 + IDC pin 7 | SPI2, 20MHz via ribbon |
| 37 | SPI_MISO | In | IDC pin 6 | SPI2 |
| 38 | AD9833_CS | Out | AD9833 FSYNC |  |
| 39 | TFT_CS | Out | IDC pin 8 |  |
| 40 | TFT_DC | Out | IDC pin 9 |  |
| 41 | TFT_RST | Out | IDC pin 10 |  |
| 42 | TOUCH_CS | Out | IDC pin 11 |  |
| 47 | RPi_TX | Out | J3 pin 3 | UART1 TX 921600 baud |
| 48 | RPi_RX | In | J3 pin 2 | UART1 RX |

# 5. UAS signal chain

1MHz acoustic chain on 3.3V rail. AD9833 sweepable 0.5–3MHz in firmware.

## 5.1 AD9833BRMZ — signal generator — CORRECTED in v2.51

CORRECTION from v2.1–v2.5: three errors removed. (1) RSET pin does not exist on the AD9833 MSOP-10 — it belongs to the AD9834. The 6.8kΩ 1% resistor and 604mVpp RSET-based amplitude calculation are removed from all tables and BOM. (2) The pin previously called "COMP" is correctly named CAP/2.5V (pin 3, per AD9833 Rev G datasheet) — this is the output of an internal 2.5V regulator used by the digital core, and requires a 100nF decoupling cap to DGND, not AGND. (3) The 200Ω external load resistor at VOUT is removed — per Rev G datasheet, an external load resistor is not required because the device has a 200Ω resistor on-board. VOUT connects directly to the AC coupling cap, then to the Tx amp.

| Parameter | Requirement |
| --- | --- |
| IC | AD9833BRMZ (MSOP-10) |
| Supply | 3.3V — both VDD pins (check module pinout for pin numbers) |
| SPI | GPIO38 (FSYNC), GPIO36 (SCLK), GPIO35 (SDATA). Write-only. SPI Mode 2 (CPOL=1, CPHA=0). |
| CAP/2.5V pin (pin 3) | 100nF X7R 0402 to DGND — internal 2.5V digital core regulator bypass. Must be connected; do not leave floating. Per Rev G datasheet: DGND, not AGND. |
| VOUT (pin 10) | Sine wave output. Amplitude approximately 0.5Vpp at 3.3V supply — fixed internally by the DAC, no external amplitude-setting resistor. Internal 200Ω output impedance — no external load resistor required or should be fitted. |
| VOUT to Tx amp | 100nF X7R 0402 AC coupling cap from VOUT directly to IC1 ch-A non-inverting input. No intermediate resistor. |
| Decoupling | 100nF X7R 0402 per VDD pin + 10µF X5R 0805 bulk |

The AD9833 output amplitude at 3.3V supply (~0.5Vpp) should be verified on the actual board with an oscilloscope before relying on the UAS attenuation measurement, since the internal amplitude is temperature- and supply-dependent. It is not adjustable in firmware.

## 5.2 Tx amplifier — IC1 ch-A, G = 4.9

NOTE: Tx amp gain was originally calculated to prevent clipping based on an incorrect 604mVpp RSET-derived amplitude. With the corrected AD9833 output of approximately 0.5Vpp: G=4.9 → output ≈2.45Vpp. This remains well within the 3.3V supply and the gain resistor values (Rf=390Ω, Rg=100Ω) are unchanged.

| Parameter | Requirement |
| --- | --- |
| IC | OPA2354 SOIC-8, ch-A of IC1. 250MHz GBW. |
| Bias | Two 10kΩ (3.3V → non-inv → GND) — 1.65V mid-rail |
| Input | 100nF AC coupling from AD9833 VOUT. Input amplitude ≈0.5Vpp. |
| Gain | Rf = 390Ω 1%, Rg = 100Ω 1%. G = 4.9. Output ≈ 2.45Vpp. Well within 3.3V supply. ✓ |
| Output | 47Ω series + 100nF coupling to J1 SMA |
| Decoupling | 100nF within 2mm V+ + 10µF bulk |

## 5.3 IC1 ch-B — unused, tied off

- Non-inv (pin 5) to 1.65V mid-rail. Inv (pin 6) to output (pin 7) — unity-gain follower.

## 5.4 Rx input protection

| Component | Value | Function |
| --- | --- | --- |
| R_in | 100Ω 0402 | Series fault current limiter |
| D1 | BAV99 SOT-23 | Clamp 0–3.3V |

## 5.5 Rx two-stage amplification — G=100 total

| Stage | Pin | Rf | Rg | Gain |
| --- | --- | --- | --- | --- |
| Stage 1 | IC2 ch-A | 910Ω 1% | 100Ω 1% | 10.1× |
| Stage 2 | IC2 ch-B | 910Ω 1% | 100Ω 1% | 10.1× — total 100×, max non-clip Rx input 32mVpp |

## 5.6 Envelope detector and ADC — GPIO1

| Component | Value | Function |
| --- | --- | --- |
| D2 | BAT54 SOD-123 | Envelope rectifier Vf≈0.2V |
| R_env | 10kΩ 0402 | RC τ=100µs, fc=1.6kHz, settling 5τ=500µs |
| C_env | 10nF X7R 0402 | RC filter |
| D3 | BAV99 SOT-23 | Hard clamp at GPIO1 |
| C_adc | 100nF X7R 0402 | GPIO1 to AGND. Within 2mm of ESP32 GPIO1 pad. Separate from C_env. |

# 6. Motor drivers — 2× BigTreeTech TMC2209 V1.x plug-in modules

U5/U6 are BTT TMC2209 V1.x plug-in modules in 2×8 sockets. Eliminates QFN-28 thermal pad failure mode; field-replaceable.

FUNCTIONAL SENSOR INPUT: StallGuard SG_RESULT feeds the PID mixing-stop algorithm. Depends on sense resistor, coil current, and SpreadCycle mode.

## 6.1 Module specifications

- Fixed onboard sense resistor: 110mΩ 1%, V1.1–V1.3 consistent. Same class as original RS1–RS4.

- Current set via onboard trimpot by default — firmware should use UART IHOLD_IRUN with I_scale_analog=0 for repeatability.

## 6.2 UART addressing

| Driver | Ref | MS1 | MS2 | Address |
| --- | --- | --- | --- | --- |
| Motor 1 | U5 | GND | GND | 0 |
| Motor 2 | U6 | 3.3V via 10kΩ | GND | 1 |

## 6.3 PDN_UART — confirmed location

VERIFIED: All BTT TMC2209 V1.1/V1.2/V1.3 modules route active UART to 4th pin down (left column), labeled PDN. 5th pin is spare — route only 4th pin, leave 5th NC.

| Component | Value | Connection |
| --- | --- | --- |
| R_UART | 1kΩ 0402 | GPIO4 → PDN junction |
| R_PDN_UP | 10kΩ 0402 | PDN junction → 3.3V |

## 6.4 SpreadCycle — firmware-enforced

Module has NO SPREAD pin. SPREAD pull-ups REMOVED from BOM. Firmware must write GCONF.en_spreadcycle=1 at every boot and read back to confirm. Load test on real hardware required before trusting SG_RESULT for PID.

## 6.5 CLK pin correction

CLK must tie to GND for internal oscillator — not float. Two 0Ω resistors (or direct traces) to GND.

## 6.6 VDD pin

3.3V assumed safe — not confirmed from module primary source. Verify before first power-on.

## 6.7 Pin mapping

| Module pin (left col) | Net | Notes |
| --- | --- | --- |
| 1. EN | GPIO7(U5)/GPIO10(U6) | 10kΩ pull-up on main board |
| 2. MS1 | GND(U5)/3.3V via 10kΩ(U6) |  |
| 3. MS2 | GND |  |
| 4. PDN (active) | UART junction | Confirmed active pad |
| 5. PDN (spare) | NC |  |
| 6. CLK | GND | Correction |
| 7. STEP | GPIO5(U5)/GPIO8(U6) |  |
| 8. DIR | GPIO6(U5)/GPIO9(U6) |  |

| Module pin (right col) | Net | Notes |
| --- | --- | --- |
| VM | 24V + bulk caps |  |
| GND | GND |  |
| A2 | J4/J5 | CAUTION: order is VM,GND,A2,A1,B1,B2,VDD,GND — A2 before A1 |
| A1 | J4/J5 |  |
| B1 | J4/J5 |  |
| B2 | J4/J5 |  |
| VDD | 3.3V | See §6.6 |
| GND | GND |  |

| Module top pins | Net |
| --- | --- |
| DIAG | NC |
| INDEX | NC |
| VREF | Onboard trimpot — no main board connection |

## 6.8 Mechanical note

Module+socket taller than QFN-28. Confirm enclosure clearance. Recommend good-retention sockets or adhesive dab given stepper motor proximity.

# 7. Limit switches

| GPIO | Ref | Connector | Wiring |
| --- | --- | --- | --- |
| GPIO14 | LIMIT_M1 | J6 | Pin 1 GND, Pin 2 GPIO14, 100nF debounce |
| GPIO15 | LIMIT_M2 | J7 | Pin 1 GND, Pin 2 GPIO15, 100nF debounce |

# 8. UI interface system — breakout board via IDC ribbon

Unchanged from v2.4/v2.5. 26-pin IDC alternative evaluated and rejected.

TFT, EC11 encoder, 2 buttons, buzzer on breakout PCB via 20-conductor IDC ribbon. Touch and display SPI tied on breakout.

## 8.1 IDC ribbon — 20-pin 2×10, odd-even

| Pin | Signal | GPIO/net | Notes |
| --- | --- | --- | --- |
| 1 | GND | GND | Red stripe ref |
| 2 | 3.3V | 3.3V |  |
| 3 | GND | GND |  |
| 4 | 5V | 5V | Buzzer |
| 5 | SPI_MOSI | GPIO35 | Shared on breakout |
| 6 | SPI_MISO | GPIO37 |  |
| 7 | SPI_CLK | GPIO36 | 20MHz max via ribbon |
| 8 | TFT_CS | GPIO39 |  |
| 9 | TFT_DC | GPIO40 |  |
| 10 | TFT_RST | GPIO41 |  |
| 11 | TOUCH_CS | GPIO42 |  |
| 12 | TOUCH_IRQ | GPIO21 | 100nF near J9 |
| 13 | EC11_A | GPIO16 |  |
| 14 | EC11_B | GPIO17 |  |
| 15 | EC11_SW | GPIO18 |  |
| 16 | BTN1 | GPIO11 |  |
| 17 | BTN2 | GPIO12 |  |
| 18 | BUZ_PWM | GPIO13 |  |
| 19 | GND | GND |  |
| 20 | NC | — | Spare |

## 8.2 TFT module — 14-pin socket on breakout

| Pin | Label | Signal |
| --- | --- | --- |
| 1 | VCC | Ribbon 2 (3.3V) |
| 2 | GND | GND |
| 3 | CS | Ribbon 8 |
| 4 | RESET | Ribbon 10 + 10kΩ pull-up |
| 5 | DC | Ribbon 9 |
| 6 | SDI | Ribbon 5 |
| 7 | SCK | Ribbon 7 |
| 8 | LED | 3.3V via 47Ω local |
| 9 | SDO | Ribbon 6 |
| 10 | T_CLK | Tied to pin 7 on breakout |
| 11 | T_CS | Ribbon 11 |
| 12 | T_DIN | Tied to pin 6 |
| 13 | T_DO | Tied to pin 9 |
| 14 | T_IRQ | Ribbon 12 |

## 8.3 EC11, buttons, buzzer

- EC11 A/B/SW → ribbon 13/14/15. BTN1/BTN2 → ribbon 16/17. Each: 10kΩ pull-up + 100nF debounce on breakout.

- Buzzer: + → ribbon 5V. − → 2N7002 drain. Gate → 10kΩ → BUZ_PWM ribbon.

# 9. External connectors (v2.51 — unchanged from v2.5)

| Ref | Label | Type | Notes |
| --- | --- | --- | --- |
| J1 | TX Transducer | SMA edge-launch | IC1 ch-A → 47Ω + 100nF → J1 |
| J2 | RX Transducer | SMA edge-launch | R_in → D1 BAV99 → IC2 |
| J3 | RPi UART | JST-XH 3-pin | Pin 1:GND / 2:RX(GPIO48) / 3:TX(GPIO47) |
| J4 | Motor 1 | JST-XH 6-pin | 1:OA1 2:NC 3:OA2 4:OB1 5:NC 6:OB2 |
| J5 | Motor 2 | JST-XH 6-pin | 1:OA1 2:NC 3:OA2 4:OB1 5:NC 6:OB2 |
| J6 | Limit M1 | JST-XH 2-pin | GND / GPIO14 |
| J7 | Limit M2 | JST-XH 2-pin | GND / GPIO15 |
| J8 | LED panel | JST-XH 2-pin | 5V / GND |
| J9 | UI breakout IDC | IDC 2×10 | 20-pin ribbon — see §8.1 |
| J13 | USB-C | USB-C receptacle | A6+B6→GPIO20, A7+B7→GPIO19 |
| J14 | PSU | 3.5mm screw terminal | Via Q1→F1→D5 |
| SK1, SK2 | TMC2209 sockets | 2×8 0.1" pin socket | Module positions U5/U6 |

# 10. PCB layout rules

## 10.1 General

- 2-layer FR4 ENIG. UAS analog top layer only, no vias in UAS paths. Antenna keepout 15mm all layers.

## 10.2 Power section

- Q1/D7 within 3mm. SS34A within 5mm of SW pin. AMS1117 tab: ≥100mm² copper + ≥4 thermal vias.

## 10.3 UAS chain

- IC1/IC2: ≥15mm separation. C_adc within 2mm of GPIO1 pad. C_env within 5mm of IC2 ch-B output.

- AD9833 VOUT connects via 100nF AC cap directly to IC1 ch-A non-inv input — no intermediate resistor.

## 10.4 TMC2209 sockets

- SK1/SK2: hand-solderable 2×8 sockets. VM bulk caps on main board near sockets. CLK GND tie within 5mm of each socket position.

## 10.5 IDC ribbon and USB-C

- J9 at board edge. 100nF at TOUCH_IRQ trace near J9.

- D+/D−: A6+B6 single net, A7+B7 single net. CC1/CC2 within 5mm.

# 11. Bill of materials — main board (v2.51)

| Qty | Ref | Description | Part | Package |
| --- | --- | --- | --- | --- |
| 1 | U1 | ESP32-S3 module | ESP32-S3-WROOM-1-N8 | Castellated |
| 1 | Q1 | P-MOSFET | AO4407A | SOIC-8 |
| 1 | D7 | Zener clamp | BZX84C15 15V | SOT-23 |
| 1 | U2 | Buck | LM2596-5.0 | TO-263-5 |
| 1 | U3 | LDO | AMS1117-3.3 | SOT-223 |
| 1 | U4 | DDS gen | AD9833BRMZ | MSOP-10 |
| 2 | IC1, IC2 | Op-amp | OPA2354 | SOIC-8 |
| 2 | SK1, SK2 | Module sockets | 2×8 0.1" pin socket | THT |
| 2 | — | TMC2209 module | BTT TMC2209 V1.x | Plug-in |
| 2 | D1, D3 | Diode clamp | BAV99 | SOT-23 |
| 1 | D2 | Rectifier | BAT54 | SOD-123 |
| 1 | D4 | Catch diode | SS34A | SMA |
| 1 | D5 | TVS | SMBJ28A | SMA |
| 1 | D6 | Schottky | 1N5819 | SOD-123 |
| 1 | F1 | Polyfuse | 1A | 1812 SMD |
| 1 | L1 | Inductor | 100µH ≥3A | SMD |
| 2 | J1, J2 | SMA | — | Board edge |
| 1 | J3 | RPi UART | JST-XH 3-pin |  |
| 2 | J4, J5 | Motors | JST-XH 6-pin |  |
| 2 | J6, J7 | Limits | JST-XH 2-pin |  |
| 1 | J8 | LED | JST-XH 2-pin |  |
| 1 | J9 | IDC | 2×10 box header |  |
| 1 | J13 | USB-C | — | SMD |
| 1 | J14 | PSU | Screw terminal |  |
| 2 | SW_boot, SW_reset | Tactile | 6×6mm SMD |  |
| 2 | LED_pwr, LED_stat | LEDs | 0805 |  |

## 11.1 Removed from BOM in v2.5 (for reference)

| Item | Reason |
| --- | --- |
| U5, U6 QFN-28 | Replaced by sockets SK1/SK2 |
| RS1–RS4 current sense | On-module |
| VCP caps (×2) | On-module |
| SPREAD pull-ups (×2) | No SPREAD pin — firmware-enforced |

## 11.2 Removed from BOM in v2.51 — errata

| Item | Reason |
| --- | --- |
| R_load 200Ω (AD9833 external load resistor) | AD9833 has internal 200Ω output impedance — no external load resistor required. Confirmed in Rev G datasheet. |
| R_RSET 6.8kΩ 1% (AD9833 amplitude setting) | RSET pin does not exist on AD9833. Belongs to AD9834. Error carried since v2.1. |

## 11.3 Breakout BOM

| Qty | Component |
| --- | --- |
| 1 | IDC 2×10 box header |
| 1 | 14-pin TFT socket |
| 1 | EC11 encoder |
| 2 | Tactile buttons |
| 1 | Passive buzzer |
| 1 | 2N7002 MOSFET |
| 7 | 10kΩ 0402 |
| 5 | 100nF 0402 |
| 1 | 20-cond IDC ribbon cable |

# 12. Pre-submission checklist (v2.51)

## 12.1 v2.51 errata checks — NEW

ERRATA CHECKS: (a) Confirm R_RSET (6.8kΩ 1%) is absent from schematic and BOM. (b) Confirm R_load (200Ω between AD9833 VOUT and AGND) is absent from schematic. (c) Confirm AD9833 VOUT connects via 100nF AC coupling cap directly to IC1 ch-A non-inv input — no intermediate resistor. (d) Confirm AD9833 CAP/2.5V pin (pin 3) connects to 100nF to DGND (not AGND). (e) Confirm ESP32 decoupling is a single 10µF+100nF pair at the single 3V3 castellated pin — not three separate decoupling groups.

## 12.2 TMC2209 module

- SK1/SK2 socket orientation confirmed. PDN: 4th pin only. CLK: GND tie fitted. EN: 10kΩ pull-up on main board.

- SPREAD pull-ups absent. RS1–RS4 absent. R_RSET and R_load absent.

VDD tolerance unverified. SpreadCycle firmware gate not yet done on real hardware.

## 12.3 USB-C

- D+ A6+B6 tied. D− A7+B7 tied. CC1/CC2 fitted. SBU NC.

## 12.4 Reverse polarity

- Q1/D7 wiring per §3.1.

## 12.5 Motor connectors

- J4/J5: 1:OA1 2:NC 3:OA2 4:OB1 5:NC 6:OB2. NC pads to GND or confirmed NC.

## 12.6 UAS and ADC

- C_adc 100nF at GPIO1 within 2mm, distinct from C_env. IC1 ch-B tied off. RSET and R_load absent.

## 12.7 IDC ribbon

- J9 pin 1 marked. 100nF at TOUCH_IRQ trace near J9.

## 12.8 Firmware notes

- TMC2209: GCONF.en_spreadcycle=1 + readback. GCONF.I_scale_analog=0 + UART IHOLD_IRUN.

- ILI9341 20MHz, XPT2046 2MHz, AD9833 SPI Mode 2. GPIO16/17: EC11 quadrature EXTI.

# 13. KiCad footprint reference (v2.51)

## 13.1 ICs and active components

| Ref | Part | KiCad footprint string |
| --- | --- | --- |
| U1 | ESP32-S3-WROOM-1-N8 | RF_Module:ESP32-S3-WROOM-1 |
| U2 | LM2596-5.0 | Package_TO_SOT_SMD:TO-263-5_TabPin3 |
| U3 | AMS1117-3.3 | Package_TO_SOT_SMD:SOT-223-3_TabPin2 |
| U4 | AD9833BRMZ | Package_SO:MSOP-10_3x3mm_P0.5mm |
| IC1, IC2 | OPA2354 | Package_SO:SOIC-8_3.9x4.9mm_P1.27mm |
| SK1, SK2 | Module socket | Connector_PinSocket_2.54mm:PinSocket_2x08_P2.54mm_Vertical |
| Q1 | AO4407A | Package_SO:SOIC-8_3.9x4.9mm_P1.27mm |

## 13.2 Diodes and protection

| Ref | Part | KiCad footprint string |
| --- | --- | --- |
| D1, D3 | BAV99 | Diode_SMD:D_SOT-23 |
| D2 | BAT54 | Diode_SMD:D_SOD-123 |
| D4 | SS34A | Diode_SMD:D_SMA |
| D5 | SMBJ28A | Diode_SMD:D_SMA |
| D6 | 1N5819 | Diode_SMD:D_SOD-123 |
| D7 | BZX84C15 | Diode_SMD:D_SOT-23 |
| F1 | Polyfuse | Fuse:Fuse_1812_4532Metric |

## 13.3 Inductor

| Ref | Part | KiCad footprint |
| --- | --- | --- |
| L1 | 100µH ≥3A | Inductor_SMD:L_Bourns-SRR1260 |

## 13.4 Resistors — CORRECTED in v2.51

CORRECTION: 200Ω (R_load, AD9833 external load) removed. 6.8kΩ precision (R_RSET, AD9833 amplitude set) removed. Both had no valid connection on the AD9833 MSOP-10.

| Group | Values — v2.51 | KiCad footprint |
| --- | --- | --- |
| Standard 0402 | 10kΩ (×16), 5.1kΩ (×2), 4.7kΩ, 1kΩ, 330Ω, 47Ω, 10Ω, 0Ω (CLK GND tie ×2) | Resistor_SMD:R_0402_1005Metric |
| Precision 1% 0402 | 390Ω (Tx Rf), 910Ω (×2 Rx Rf), 100Ω (×3 Rg) | Resistor_SMD:R_0402_1005Metric |

200Ω and 6.8kΩ removed in v2.51. 2512 current sense also removed in v2.5 (on TMC2209 module).

## 13.5 Capacitors — ceramic 0402 — CORRECTED in v2.51

CORRECTION: 100nF count reduced from 25 to 23 (removed two extra ESP32 decoupling caps — module has one 3V3 pin, not three). 10µF 0805 count reduced from 5 to 4 (one less ESP32 bulk cap).

| Value | Footprint | Qty / notes |
| --- | --- | --- |
| 100nF X7R 0402 | Capacitor_SMD:C_0402_1005Metric | 23 total (was 25 in v2.5) — decoupling, debounce, AC coupling, C_adc, CAP/2.5V pin |
| 10nF X7R 0402 | Capacitor_SMD:C_0402_1005Metric | 2 — C_env + AD9833 CAP/2.5V is separately a 100nF (above) |
| 1nF X7R 0402 | Capacitor_SMD:C_0402_1005Metric | 1 — SW snubber |
| 1µF X7R 0402 | Capacitor_SMD:C_0402_1005Metric | 1 — C_en (ESP32 EN RC) |

## 13.6 Capacitors — 0805, tantalum, electrolytic

| Value/type | KiCad footprint | Qty |
| --- | --- | --- |
| 10µF X5R 0805 (CORRECTED: 4 not 5) | Capacitor_SMD:C_0805_2012Metric | 4 — ESP32 (×1), AD9833, IC1, IC2 |
| 10µF 10V tantalum Case A | Capacitor_Tantalum_SMD:CP_EIA-3216-18_Kemet-A | 2 — AMS1117 in/out |
| 100µF 50V radial | Capacitor_THT:C_Radial_D8.0mm_P3.5mm | 1 — LM2596 input |
| 220µF 10V low-ESR radial | Capacitor_THT:C_Radial_D8.0mm_P3.5mm | 1 — LM2596 output |
| 100µF 35V radial | Capacitor_THT:C_Radial_D6.3mm_P2.5mm | 2 — TMC2209 VM bulk per socket |

## 13.7 Connectors — main board

| Ref | Description | KiCad footprint |
| --- | --- | --- |
| J1, J2 | SMA edge-launch | Connector_Coaxial:SMA_Molex_0732511150_EdgeMount |
| J3 | JST-XH 3-pin | Connector_JST:JST_XH_B3B-XH-A_1x03_P2.50mm_Vertical |
| J4, J5 | JST-XH 6-pin | Connector_JST:JST_XH_B6B-XH-A_1x06_P2.50mm_Vertical |
| J6, J7 | JST-XH 2-pin | Connector_JST:JST_XH_B2B-XH-A_1x02_P2.50mm_Vertical |
| J8 | JST-XH 2-pin | Connector_JST:JST_XH_B2B-XH-A_1x02_P2.50mm_Vertical |
| J9 | IDC 2×10 | Connector_IDC:IDC-Header_2x10_P2.54mm_Vertical |
| J13 | USB-C | Connector_USB:USB_C_Receptacle_HRO_TYPE-C-31-M-12 |
| J14 | Screw terminal | TerminalBlock_Phoenix:TerminalBlock_Phoenix_PT-1,5-2-3.5_1x02_P3.50mm_Horizontal |
| SK1, SK2 | Module sockets | Connector_PinSocket_2.54mm:PinSocket_2x08_P2.54mm_Vertical |

## 13.8 Breakout footprints

| Component | KiCad footprint |
| --- | --- |
| IDC 2×10 (breakout) | Connector_IDC:IDC-Header_2x10_P2.54mm_Vertical |
| 14-pin TFT socket | Connector_PinHeader_2.54mm:PinSocket_1x14_P2.54mm_Vertical |
| EC11 encoder | Rotary_Encoder:RotaryEncoder_Alps_EC11E-Switch_Vertical_H20mm |
| 6×6mm tactile THT | Button_Switch_THT:SW_Tactile_Straight_KSA0Axx1LFTR |
| 2N7002 SOT-23 | Package_TO_SOT_SMD:SOT-23 |

## 13.9 Switches and LEDs

| Ref | Part | KiCad footprint |
| --- | --- | --- |
| SW_boot, SW_reset | 6×6mm SMD tactile | Button_Switch_SMD:SW_SPST_EVQP2_6x6mm |
| LED_pwr | Red LED 0805 | LED_SMD:LED_0805_2012Metric |
| LED_stat | Green LED 0805 | LED_SMD:LED_0805_2012Metric |

— End of document —